Annex No. 1 – Technical specifications

Technical specifications

Electric Impedance Tomography

- The subject of the public contract is the supply of measuring equipment for recording 16-channel electrical impedance tomography (EIT).

The subject of performance also includes transport to the place of performance.

Manufacturer of the device:				to be completed by the participant

Exact type designation of the device:		to be completed by the participant

Electrical impedance tomography must meet following criterias:

| Basic technical parameters | Minimal required values – obligatory condition! | Value of the offered device |
| --- | --- | --- |
| Excitation and Measurement Frequency |  |  |
| Range | min. 100 Hz to 1 MHz | the participiant indicates YES/NO and indicates specific value |
| Resolution | min. 40μHz @ f< 10kHz min. 2mHz @ 10kHz < f <500kHz min. 220mHz @ f>500kHz | the participiant indicates YES/NO and indicates specific value |
| Precision Absolute | min. ± 100 ppm (at 25 °C) | the participiant indicates YES/NO and indicates specific value |
| Temperature Drift | max. ± 10 ppm over operating temperature range | the participiant indicates YES/NO and indicates specific value |
| Excitation signal |  |  |
| Range | min. 100 nA to 10 mA peak-amplitude | the participiant indicates YES/NO and indicates specific value |
| Resolution | max. 3 μA | the participiant indicates YES/NO and indicates specific value |
| Maximum Continuous Output Current | min. 50 mA | the participiant indicates YES/NO and indicates specific value |
| Compliance Voltage | max. ± 11 V (battery powered ± 5 V) | the participiant indicates YES/NO and indicates specific value |
| Measurement Timing |  |  |
| Frame Rate: 0.1 ... 100 fps (refer to Functional Description for information on constraints towards the maximum allowable frame rate) | Yes | the participiant indicates YES/NO |
| Absolute Time Base Precision | min. ± 100 ppm | the participiant indicates YES/NO and indicates specific value |
| Frequency Sweep Settings |  |  |
| Sweep type: linear, logarithmic, list | Yes | the participiant indicates YES/NO |
| Points | min. 1 to 128 | the participiant indicates YES/NO and indicates specific value |
| Measurement Terminal Specifications |  |  |
| Number of dual role ports with high input impedance voltage measurement and optional current injection function | at least 16 | the participiant indicates YES/NO and indicates specific value |
| Input Voltage Range | ±10 V (battery powered ±5V) | the participiant indicates YES/NO and indicates specific value |
| Input Impedance (in non-injecting state) | min. 1 TΩ, 2,3 pF | the participiant indicates YES/NO and indicates specific value |
| Absolute Max. Short Circuit Current (Source/Sink) For Any Port In Injection State | min. 120 mA | the participiant indicates YES/NO and indicates specific value |
| Measurement Port Shielding: Grounded passive (capacitive) shielding is available | Yes | the participiant indicates YES/NO |
| Ethernet Interface |  |  |
| Standard Conformity: 10/100 Base-T, RJ45 (TCP/IP) | Yes | the participiant indicates YES/NO |
| USB Interface |  |  |
| Standard Conformity: USB 2.0, FS (12 Mbits/s) | Yes | the participiant indicates YES/NO |
| Connector: Micro USB type B (full speed USB) | Yes | the participiant indicates YES/NO |
| Isolation: 2500 V rms for 1 minute per UL 1577 IEC 60950-1: 600 V rms (basic) DIN V VDE V 0884-10 (VDE V 0884-10):2006-12 VIORM = 560 V peak | Yes | the participiant indicates YES/NO |
| General Specifications |  |  |
| Power Requirements: 12VDC (typ.), 18W (max.) | Yes | the participiant indicates YES/NO |
| Weight | max. 1,2 kg | the participiant indicates YES/NO and indicates specific value |
| Operating Conditions: 0°C to 40°C, <80% relative humidity non condensing, 0...3000 m altitude | Yes | the participiant indicates YES/NO |

The participiant should enter the data in compliance with the technical specification of the offered devices.